## Slide 1
Inspiratory Arm
Expiratory Arm
Ventilator
Inspiratory Splitter Box
Expiratory Splitter Box
A
B
Filter
Filter
Patient A Tubing
Patient B Tubing
A
B
Filter
Filter
Patient A Tubing
Patient B Tubing
This design will allow a modern ventilator to have a splitter. The key innovation here came from the Differential Multiventilation group– the Bias Circuit. That circuit bypasses the compensatory electronics inside the ventilator by providing the ventilator with immediate feedback for the bias flow Inspiratory and Expiratory Splitter Boxes are on the pages to follow. #project-4-way-ventilator, Helpful Engineering (rev 9, 20 May 2020)
Bias Circuit

## Slide 2
Ventilator Splitter Box (VSB)
Inspiratory Section
Inspiratory
Arm
Expiratory
Arm
Ventilator
Expiratory Section
P1 Shutoff
P1 Flow Meter
P2 Shutoff
P2 Flow Meter
P3 Shutoff
P3 Flow Meter
P4 Shutoff
P4 Flow Meter
P1 Inspiration
P2 Inspiration
P3 Inspiration
P4 Inspiration
P1 Expiration
P2 Expiration
P3 Expiration
P4 Expiration
Tetra Schematics v11
2 June 2020
Bias Circuit
1-Way Valve
1-Way Valve
1-Way Valve
1-Way Valve
1-Way Valve
1-Way Valve
P1 Shutoff
P1 Flow Control
P2 Shutoff
P2 Flow Control
P3 Shutoff
P3 Flow Control
P4 Shutoff
P4 Flow Control
1-Way Valve
1-Way Valve
1-Way Valve
1-Way Valve

## Slide 3
Spring
Flow
Cap
Seal
Airflow
Airflow